## APMS QExactive MS
Method of Q Exactive Plus
OVERALL METHOD SETTINGS
Global Settings
use lock masses best
Lock mass injection ―
Chrom. peak width (FWHM) 15 s
Time
Method duration 75.00 min
Customized Tolerances (+/-)
Lock Masses ―
Inclusion ―
Exclusion ―
Neutral Loss ―
Mass Tags ―
Dynamic Exclusion ―
Experiment
FULL MS / DD-MS² (TOPN)
General
Runtime 0 to 75 min
Polarity Positive
In-source CID 0.0 eV
Default charge state 2
Inclusion ―
Exclusion ―
Tags ―
Full MS
Microscans 1
Resolution 70 | 0
AGC target 1e6
Maximum IT 250 ms
Number of scan ranges 1
Scan range 300 to 1500 m/z
Spectrum data type Profile
dd-MS² / dd-SIM
Microscans 1
Resolution 17 | 500
AGC target 5e4
Maximum IT 60 ms
Loop count 20
MSX count 1
TopN 20
Isolation window 2.0 m/z
Isolation offset 0.0 m/z
Scan range 200 to 2000 m/z
Fixed first mass ―
NCE / stepped NCE 26
Spectrum data type Centroid
dd Settings
Minimum AGC target 8.00e3
Intensity threshold 1.3e5
Apex trigger ―
Charge exclusion unassigned | 1
Peptide match Preferred
Exclude isotopes on
Dynamic exclusion auto
If idle .. do not pick others
Setup
TUNEFILES
General
Switch Count 0
Base Tunefile C:\Xcalibur\methods\BEST_TUNE.mstune
CONTACT CLOSURE
General
Used False
Start in Closed True
Switch Count 0
SYRINGE
General
Used False
Start in OFF True
Stop at end of run False
Switch Count 0
Pump setup
Syringe type Hamilton
Flow rate 3.000 µL/min
Inner diameter 2.303 mm
Volume 250 µL
DIVERT VALVE A
General
Used False
Start in 1-2 True
Switch Count 0
DIVERT VALVE B
General
Used False
Start in 1-2 True
Switch Count 0
LOCK MASSES
3 entries
Mass Polarity Start End Comment
[m/z] [min] [min]
371.10124 Positive
445.12003 Positive
519.13883 Positive
INCLUSION LIST
(no entries)
EXCLUSION LIST
(no entries)
NEUTRAL LOSSES
(no entries)
MASS TAGS
(no entries)

## APMS QExactive HPLC
n1200 Easy LC
Volume [µl] : 10.00
Flow [µl / min] : 20.00
Sample loading:
Volume [µl] : 6.00
Flow [µl / min] : 1.00
Max. pressure [Bar] : 800.00
Gradient:
Time [mm:ss] Duration [mm:ss] Flow [nl/min] Mixture [%B]
00:00 00:00 300 2
01:00 01:00 300 4
54:00 53:00 300 32
61:00 07:00 300 85
71:00 10:00 300 85
72:00 01:00 300 2
75:00 03:00 300 2
Pre-column equilibration:
Volume [µl] : 5.00
Flow [µl / min] : (unspecified)
Max. pressure [Bar] : 800.00
Analytical column equilibration:
Volume [µl] : 6.00
Flow [µl / min] : (unspecified)
Max. pressure [Bar] : 800.00
Autosampler wash:
Flush volume [µl] : 100.00

## APMS Orbitrap Elite MS
Method of Oribtrap Elite
Creator: Thermo Scientific
MS Run Time (min): 90.00
Sequence override of method parameters not enabled.
Divert Valve: not used during run
Contact Closure: not used during run
Syringe Pump: not used during run
MS Detector Settings:
Real-time modifications to method disabled
Stepped collision energy not enabled
Additional Microscans:
MS2 0 0
MS3 0 0
MS4 0 0
MS5 0 0
MS6 0 0
MS7 0 0
MS8 0 0
MS9 0 0
MS10 0 0
Experiment Type: Nth Order Double Play
Tune Method: 092613-NSI-VE-injection-times
Scan Event Details:
1: FTMS + p norm o(300.0-2000.0)
CV = 0.0V
2: ITMS + c norm Dep Rapid MS/MS Most intense ion from (1)
Activation Type: CID
Min. Signal Required: 500.0
Isolation Width: 2.00
Normalized Coll. Energy: 35.0
Default Charge State: 2
Activation Q: 0.250
Activation Time: 10.000
CV = 0.0V
Scan Event 2 repeated for top 20 peaks.
Lock Masses:
Pos List Name: N/A
Source: API Source
Mass List: (none)
Neg List Name: N/A
Source: API Source
Mass List: (none)
Data Dependent Settings:
Use separate polarity settings disabled
Parent Mass List: (none)
Reject Mass List: (none)
Neutral Loss Mass List: (none)
Product Mass List: (none)
Neutral loss in top: 3
Product in top: 3
Most intense if no parent masses found not enabled
Add/subtract mass not enabled
FT master scan preview mode enabled
Charge state screening enabled
Charge state dependent ETD time not enabled
Monoisotopic precursor selection enabled
Charge state rejection enabled
Unassigned charge states : rejected
Charge state 1 : rejected
Charge state 2 : not rejected
Charge state 3 : not rejected
Charge states 4+ : not rejected
Chromatography mode is disabled
Global Data Dependent Settings:
Predict ion injection time enabled
Use global parent and reject mass lists not enabled
Exclude parent mass from data dependent selection not enabled
Exclusion mass width relative to mass
Exclusion mass width relative to low (ppm): 10.00
Exclusion mass width relative to high (ppm): 10.00
Parent mass width by mass
Parent mass width low: 0.5000
Parent mass width high: 0.5000
Reject mass width relative to mass
Reject mass width relative to low (ppm): 10.00
Reject mass width relative to high (ppm): 10.00
Zoom/UltraZoom scan mass width by mass
Zoom/UltraZoom scan mass width low: 5.00
Zoom/UltraZoom scan mass width high: 5.00
FT SIM scan mass width low: 5.00
FT SIM scan mass width high: 5.00
Neutral Loss candidates processed by decreasing intensity
Neutral Loss mass width by mass
Neutral Loss mass width low: 0.5000
Neutral Loss mass width high: 0.5000
Product candidates processed by decreasing intensity
Product mass width by mass
Product mass width low: 0.5000
Product mass width high: 0.5000
MS mass range: 0.00-1000000.00
MSn mass range by mass
MSn mass range: 0.00-1000000.00
Use m/z values as masses not enabled
Analog UV data dep. not enabled
Dynamic exclusion enabled
Repeat Count: 1
Repeat Duration: 40.00
Exclusion List Size: 500
Exclusion Duration: 40.00
Exclusion mass width relative to mass
Exclusion mass width relative to low (ppm): 10.00
Exclusion mass width relative to high (ppm): 10.00
Expiration: disabled
Isotopic data dependence not enabled
Custom Data Dependent Settings:
Not enabled

## APMS Oribtrap Elite HPLC
n1200 Easy LC
Sample pickup:
Volume [µl] : 3.00
Flow [µl / min] : 20.00
Sample loading:
Volume [µl] : 7.00
Flow [µl / min] : (unspecified)
Max. pressure [Bar] : 850.00
Gradient:
Time [mm:ss] Duration [mm:ss] Flow [nl/min] Mixture [%B]
00:00 00:00 300 2
01:00 01:00 300 5
72:00 71:00 300 30
73:00 01:00 300 95
79:00 06:00 300 95
80:00 01:00 400 2
90:00 10:00 400 2
Pre-column equilibration:
Volume [µl] : 0.00
Flow [µl / min] : (unspecified)
Max. pressure [Bar] : 250.00
Analytical column equilibration:
Volume [µl] : 6.00
Flow [µl / min] : (unspecified)
Max. pressure [Bar] : 850.00
Auto-sampler wash:
Flush volume [µl] : 100.00

## APMS MaxQuant Search parameters
Parameter | Value
Version | 1.6.2.10
Include contaminants | TRUE
PSM FDR | 0.01
PSM FDR Crosslink | 0.01
Protein FDR | 0.01
Site FDR | 0.01
Use Normalized Ratios For Occupancy | TRUE
Min. peptide Length | 7
Min. score for unmodified peptides | 0
Min. score for modified peptides | 40
Min. delta score for unmodified peptides | 0
Min. delta score for modified peptides | 6
Min. unique peptides | 0
Min. razor peptides | 1
Min. peptides | 1
Use only unmodified peptides and | TRUE
Modifications included in protein quantification | Oxidation (M);Acetyl (Protein N-term)
Peptides used for protein quantification | Razor
Discard unmodified counterpart peptides | TRUE
Label min. ratio count | 2
Use delta score | FALSE
iBAQ | FALSE
iBAQ log fit | FALSE
Match between runs | TRUE
Matching time window [min] | 1
Alignment time window [min] | 20
Find dependent peptides | FALSE
Fasta file | C:\MaxQuant\MaxQuant\FASTA Files\uniprot-reviewed-human-canonical-20180321.fasta
Decoy mode | revert
Include contaminants | TRUE
Advanced ratios | TRUE
Fixed andromeda index folder
Temporary folder
Combined folder location
Second peptides | TRUE
Stabilize large LFQ ratios | TRUE
Separate LFQ in parameter groups | FALSE
Require MS/MS for LFQ comparisons | TRUE
Calculate peak properties | FALSE
Main search max. combinations | 200
Advanced site intensities | TRUE
Write msScans table | TRUE
Write msmsScans table | TRUE
Write ms3Scans table | TRUE
Write allPeptides table | TRUE
Write mzRange table | TRUE
Write pasefMsmsScans table | TRUE
Write accumulatedPasefMsmsScans table | TRUE
Max. peptide mass [Da] | 4600
Min. peptide length for unspecific search | 8
Max. peptide length for unspecific search | 25
Razor protein FDR | TRUE
Disable MD5 | FALSE
Max mods in site table | 3
Match unidentified features | FALSE
Epsilon score for mutations
Evaluate variant peptides separately | TRUE
Variation mode | None
MS/MS tol. (FTMS) | 20 ppm
Top MS/MS peaks per Da interval. (FTMS) | 12
Da interval. (FTMS) | 100
MS/MS deisotoping (FTMS) | TRUE
MS/MS deisotoping tolerance (FTMS) | 7
MS/MS deisotoping tolerance unit (FTMS) | ppm
MS/MS higher charges (FTMS) | TRUE
MS/MS water loss (FTMS) | TRUE
MS/MS ammonia loss (FTMS) | TRUE
MS/MS dependent losses (FTMS) | TRUE
MS/MS recalibration (FTMS) | FALSE
MS/MS tol. (ITMS) | 0.5 Da
Top MS/MS peaks per Da interval. (ITMS) | 8
Da interval. (ITMS) | 100
MS/MS deisotoping (ITMS) | FALSE
MS/MS deisotoping tolerance (ITMS) | 0.15
MS/MS deisotoping tolerance unit (ITMS) | Da
MS/MS higher charges (ITMS) | TRUE
MS/MS water loss (ITMS) | TRUE
MS/MS ammonia loss (ITMS) | TRUE
MS/MS dependent losses (ITMS) | TRUE
MS/MS recalibration (ITMS) | FALSE
MS/MS tol. (TOF) | 40 ppm
Top MS/MS peaks per Da interval. (TOF) | 10
Da interval. (TOF) | 100
MS/MS deisotoping (TOF) | TRUE
MS/MS deisotoping tolerance (TOF) | 0.01
MS/MS deisotoping tolerance unit (TOF) | Da
MS/MS higher charges (TOF) | TRUE
MS/MS water loss (TOF) | TRUE
MS/MS ammonia loss (TOF) | TRUE
MS/MS dependent losses (TOF) | TRUE
MS/MS recalibration (TOF) | FALSE
MS/MS tol. (Unknown) | 0.5 Da
Top MS/MS peaks per Da interval. (Unknown) | 8
Da interval. (Unknown) | 100
MS/MS deisotoping (Unknown) | FALSE
MS/MS deisotoping tolerance (Unknown) | 0.15
MS/MS deisotoping tolerance unit (Unknown) | Da
MS/MS higher charges (Unknown) | TRUE
MS/MS water loss (Unknown) | TRUE
MS/MS ammonia loss (Unknown) | TRUE
MS/MS dependent losses (Unknown) | TRUE
MS/MS recalibration (Unknown) | FALSE
Site tables | Oxidation (M)Sites.txt

## Parental cell MaxQuant params
Parameter | Value
Version | 1.5.8.2
Fixed modifications | Carbamidomethyl (C)
Include contaminants | TRUE
PSM FDR | 0.01
Protein FDR | 0.01
Site FDR | 0.01
Use Normalized Ratios For Occupancy | TRUE
Min. peptide Length | 7
Min. score for unmodified peptides | 0
Min. score for modified peptides | 40
Min. delta score for unmodified peptides | 0
Min. delta score for modified peptides | 6
Min. unique peptides | 0
Min. razor peptides | 1
Min. peptides | 1
Use only unmodified peptides and | TRUE
Modifications included in protein quantification | Oxidation (M);Acetyl (Protein N-term)
Peptides used for protein quantification | Razor
Discard unmodified counterpart peptides | TRUE
Label min. ratio count | 2
Use delta score | FALSE
iBAQ | FALSE
iBAQ log fit | FALSE
Match between runs | TRUE
Matching time window [min] | 1.5
Alignment time window [min] | 20
Find dependent peptides | FALSE
Fasta file | C:\MaxQuant\MaxQuant\FASTA Files\uniprot_canonical.fasta
Decoy mode | revert
Include contaminants | TRUE
Advanced ratios | TRUE
Fixed andromeda index folder
Temporary folder
Combined folder location
Second peptides | TRUE
Stabilize large LFQ ratios | TRUE
Separate LFQ in parameter groups | FALSE
Require MS/MS for LFQ comparisons | TRUE
Calculate peak properties | FALSE
Main search max. combinations | 200
Advanced site intensities | TRUE
Write msScans table | TRUE
Write msmsScans table | TRUE
Write ms3Scans table | TRUE
Write allPeptides table | TRUE
Write mzRange table | TRUE
Max. peptide mass [Da] | 4600
Min. peptide length for unspecific search | 8
Max. peptide length for unspecific search | 25
Razor protein FDR | TRUE
Disable MD5 | FALSE
Max mods in site table | 3
Match unidentified features | FALSE
MS/MS tol. (FTMS) | 20 ppm
Top MS/MS peaks per Da interval. (FTMS) | 12
Da interval. (FTMS) | 100
MS/MS deisotoping (FTMS) | TRUE
MS/MS deisotoping tolerance (FTMS) | 7
MS/MS deisotoping tolerance unit (FTMS) | ppm
MS/MS higher charges (FTMS) | TRUE
MS/MS water loss (FTMS) | TRUE
MS/MS ammonia loss (FTMS) | TRUE
MS/MS dependent losses (FTMS) | TRUE
MS/MS recalibration (FTMS) | FALSE
MS/MS tol. (ITMS) | 0.5 Da
Top MS/MS peaks per Da interval. (ITMS) | 8
Da interval. (ITMS) | 100
MS/MS deisotoping (ITMS) | FALSE
MS/MS deisotoping tolerance (ITMS) | 0.15
MS/MS deisotoping tolerance unit (ITMS) | Da
MS/MS higher charges (ITMS) | TRUE
MS/MS water loss (ITMS) | TRUE
MS/MS ammonia loss (ITMS) | TRUE
MS/MS dependent losses (ITMS) | TRUE
MS/MS recalibration (ITMS) | FALSE
MS/MS tol. (TOF) | 40 ppm
Top MS/MS peaks per Da interval. (TOF) | 10
Da interval. (TOF) | 100
MS/MS deisotoping (TOF) | TRUE
MS/MS deisotoping tolerance (TOF) | 0.01
MS/MS deisotoping tolerance unit (TOF) | Da
MS/MS higher charges (TOF) | TRUE
MS/MS water loss (TOF) | TRUE
MS/MS ammonia loss (TOF) | TRUE
MS/MS dependent losses (TOF) | TRUE
MS/MS recalibration (TOF) | FALSE
MS/MS tol. (Unknown) | 0.5 Da
Top MS/MS peaks per Da interval. (Unknown) | 8
Da interval. (Unknown) | 100
MS/MS deisotoping (Unknown) | FALSE
MS/MS deisotoping tolerance (Unknown) | 0.15
MS/MS deisotoping tolerance unit (Unknown) | Da
MS/MS higher charges (Unknown) | TRUE
MS/MS water loss (Unknown) | TRUE
MS/MS ammonia loss (Unknown) | TRUE
MS/MS dependent losses (Unknown) | TRUE
MS/MS recalibration (Unknown) | FALSE
Site tables | Oxidation (M)Sites.txt